| Project | PIC iX C.0x PIC iX 4.x | DHF Number | 453564736341 |
| --- | --- | --- | --- |

| Revision | Change Description in Brief |
| --- | --- |
| DV | Update for May patch qualifications. Corrected .NET line items in the Windows Server 2019 and Windows Server 2022 Sections to mention ‘Server’ rather than ‘Client’. |

- Reviews and approvals are performed according to A-Q2920-02237 Design and Development Planning.

Contents

1	Purpose	3

2	Scope	3

3	Terminology & Abbreviations	3

4	Document creation work instructions	3

5	References	3

6	Document Revision History	3

7	Document Information	5

7.1	Document Properties	5

8	Appendices	5

8.1	Windows 10 Enterprise LTSB v1607	6

8.2	Windows Server 2016	36

8.3	Windows 10 Enterprise LTSC v1809	65

8.4	Windows Server 2019	102

8.5	Windows 10 IoT Enterprise LTSC v21H2	137

8.6	Windows Server 2022	153

8.7	Validated AVS	169

8.8	Hypervisor Compatibility	173

# Purpose

This document is used to provide customers and the field organization with a single source for product security advisory that is not already present in the Manufacturer Disclosure Statement for Medical Device Security (MDS2) form. It describes the status of known product or service related security vulnerabilities and patches and is intended to be maintained throughout the product lifecycle, while the security features listed in the MDS2 are stable because it is related to the product design.

# Scope

This document is applicable to PIC iX C.0x and 4.x inclusive of PIC iX Servers and Client Platforms: Windows 10 Enterprise LTSB v1607, Windows Server 2016, Windows 10 Enterprise LTSC v1809, and Windows Server 2019, Windows 10 IoT Enterprise LTSC v21H2, Windows Server 2022, as well as third-party applications.

# Terminology & Abbreviations

| Terminology & Abbreviations | Description/Definition |
| --- | --- |
| N/A | N/A |

# Document creation work instructions

The derived customer-facing document shall not contain the headers and footers of the approved document. Only provide the information after the section break in the "Appendices" chapter to the customer.

Please ensure that the derived customer facing document is also reviewed by your local Legal and Regulatory Representative before this is made available to (the) customer(s). Especially the specific statements for Scoring: and Recommended Customer Action: need attention.

# References

| Reference Number | Document Title | Document ID |
| --- | --- | --- |
| 1 | Philips Patient Monitoring System Security for Clinical Networks | 453564943301 |

# Document Revision History

| Revision | Release Date | Author | Description of changes | CR / Reason |
| --- | --- | --- | --- | --- |
| DU | 21-Apr-2026 | Atia Wali | Update for April patch qualifications. | EC 2917-2025-12-29900 |
| DT | 26-Mar-2026 | Atia Wali | Update for PIC iX 4.7 release. | EC 2917-2025-11-29810 |
| DR | 20-Mar-2026 | Atia Wali | Update for March patch qualifications. | EC-2917-2025-12-29899 |
| DP | 23-Feb-2026 | Atia Wali | Update for February patch qualifications. | EC-2917-2025-12-29898 |
| DN | 09-Feb-2026 | Atia Wali | Update for PIC iX 4.6 release. | EC-2917-2025-10-29767 |
| DM | 29-Jan-2026 | Atia Wali | Update for January patch qualifications. | EC-2917-2025-12-29897 |
| DL | 17-Dec-2025 | Atia Wali | Update for December patch qualifications. | EC-2917-2025-06-29411 |
| DK | 25-Nov-2025 | Atia Wali | Update for November patch qualifications. | EC-2917-2025-06-29410 |
| DJ | 12-Nov-2025 | Atia Wali | Update for Critical Vulnerability patch qualifications. | EC-2917-2025-11-29778 |
| DH | 22-Oct-2025 | Atia Wali | Update for October patch qualifications. | EC-2917-2025-06-29409 |
| DG | 17-Oct-2025 | Atia Wali | Update for PIC iX 4.5 Release. | EC-2917-2024-10-28782 |
| DF | 17-Sep-2025 | Atia Wali | Update for September patch qualifications. New column added to indicate Microsoft Product for each update. Corrected Patch ID in Section 8.1 and 8.2 for August patch qualification for Security Update for SQL Server 2016 Service Pack 3 GDR. Updated Section 8.7 to reflect PIC iX 4.4 qualifications. | EC-2917-2025-06-29408 |
| DE | 20-Aug-2025 | Atia Wali | Update for August patch qualifications. | EC-2917-2025-06-29407 |
| DD | 23-Jul-2025 | Atia Wali | Update for July patch qualifications. | EC-2917-2025-06-29406 |
| DC | 19-Jun-2025 | Atia Wali | Update for June patch qualifications. | EC-2917-2025-01-28972 |
| DB | 22-May-2025 | Atia Wali | Update for May patch qualifications. | EC-2917-2025-01-28971 |
| DA | 24-Apr-2025 | Atia Wali | Update for April patch qualifications. Corrected and updated Section 8.7, including support for CrowdStrike 7.23. | EC-2917-2025-01-28970 |
| CZ | 28-Mar-2025 | Atia Wali | Update for March patch qualifications. Corrected and updated Section 8.7, including support for Windows Defender for Windows 10 v21H2 and Windows Server 2022. | EC-2917-2025-01-28969 |
| CY | 03-Mar-2025 | Atia Wali | Update for February patch qualifications. Updated Section 2 to include all Operating Systems. | EC-2917-2025-01-28968 |
| CW | 03-Feb-2025 | Atia Wali | Update for January patch qualifications. Corrected and updated table in Section 8.7, including clarification on support for Trend Micro 20.0.1 and Crowdstrike 7.19. | EC 2917-2025-01-28967 |
| CV | 19-Dec-2024 | Atia Wali | Update for December patch qualifications | EC 2917-2024-07-28558 |
| CU | 25-Nov-2024 | Atia Wali | Update for November patch qualifications | EC 2917-2024-07-28557 |
| CT | 18-Nov-2024 | Atia Wali | PIC iX 4.3.3 Release: - Added support for Hyper-V Server 2022 in Section 8.6 | EC 2917-2024-10-28831 |
| CR | 18-Oct-2024 | Atia Wali | Update for October patch qualifications. Added KB5040946 to Windows 10 IoT Enterprise section. | EC 2917-2024-07-28556 |
| CP | 01-Oct-2024 | Atia Wali | Update for September patch qualifications | EC 2917-2024-07-28555 |
| CN | 06-Sep-2024 | Atia Wali | Update for August patch qualifications Corrected patch ID number for July patch qualification for dotnet-hosting-6.0.32-winows installation. | EC 2917-2024-07-28554 |
| CM | 04-Aug-2024 | Atia Wali | Update for July patch qualifications | EC 2917-2024-07-28553 |
| CL | 08-Jul-2024 | Atia Wali | Update for June patch qualifications | EC 2917-2024-01-28105 |
| CK | 17-Jun-2024 | Atia Wali | Update for May patch qualifications | EC 2917-2024-01-28104 |
| CJ | 10-May-2024 | Atia Wali | Update for April patch qualifications | EC 2917-2024-01-28103 |

# Document Information

## Document Properties

| Property | Value |
| --- | --- |
| Template Used | A-Q2920-01684-T011 |
| Document Class | See system managing this document |
| Tier | See system managing this document |
| Element | See system managing this document |

# Appendices

Product security advisory

Philips continuously assesses vulnerability notifications and security updates of OS and other components as documented in the ‘Philips and cybersecurity position paper’. This product specific security advisory informs you about vulnerabilities and security updates that may require your attention or action.

For additional security information such as the ‘Philips and cybersecurity position paper’, the Manufacturer Disclosure Statement for Medical Device Security (MDS2) and general security advisories, please visit the Philips product security internet site at https://www.philips.com/security.

Product vulnerabilities

The table below is to inform you about the vulnerabilities that are either not mitigated effectively or where there is a need to communicate about specifically.

| Product Vulnerability | Evaluation Date | Scoring (<Scoring system, for example CVSS 3.1>) | Applicable for Release | Recommended Customer Action |
| --- | --- | --- | --- | --- |
| N/A | N/A | N/A | N/A | N/A |

Validated patches

Below table is to inform you about the Philips validated patches that can be applied in accordance with the patch policy defined in the Philips Patient Monitoring System Security for Clinical Networks document.

Note: Vulnerability/Patch ID marked with “*” indicates it is optional.

## Windows 10 Enterprise LTSB v1607

| Release: | PIC iX C.0x | OPERATING_SYSTEM / SERVICE_PACK: (The Main Operating System is the OS of the system component that connects to the network) | PIC iX Servers and Client Application: Windows 10 Enterprise LTSB v1607 |
| --- | --- | --- | --- |

| Windows 10 Enterprise LTSB v1607 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Vulnerability / Patch ID | Last Update by Philips | Activity Status | Recommended Customer Action | Notes / Instructions | Microsoft Product |
| KB5088064 | 27-May-2026 | Solution Available | Install Recommended Update prior to other updates | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5087537 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5087065 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 27-May-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.141 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077473** | 27-May-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 CU | Microsoft SQL Server 2016 |
| KB5082089 | 21-Apr-2026 | Solution Available | Superseded by KB5088064 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| **Only for customers using the Azure Connect Feature update path. KB5082198 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5082411 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 21-Apr-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Apr-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.140 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077474 | 21-Apr-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5078938 | 20-Mar-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5075902 | 23-Feb-2026 | Solution Available | Superseded by KB5082089 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5075999 | 23-Feb-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 23-Feb-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 23-Feb-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.139 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5073447 | 26-Jan-2026 | Solution Available | Superseded by KB5075902 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5073722 | 26-Jan-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5068401 | 26-Jan-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5071543 | 17-Dec-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 17-Dec-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 17-Dec-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool - v5.138 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5070247 | 25-Nov-2025 | Solution Available | Superseded by KB5073447 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5068864 | 25-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 25-Nov-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool - v5.137 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5066584 | 22-Oct-2025 | Solution Available | Superseded by KB5070247 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5066836 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5066136 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 22-Oct-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Oct-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool - v5.136 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5065687 | 17-Sep-2025 | Solution Available | Superseded by KB5066584 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5065427 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5065749 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5063871 | 20-Aug-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 20-Aug-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.135 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 20-Aug-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5058718 | 20-Aug-2025 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5062799 | 20-Jul-2025 | Solution Available | Superseded by KB5065687 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5062560 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5062064 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5060954 | 19-Jun-2025 | Solution Available | Superseded by KB5062799 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5061010 | 19-Jun-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 19-Jun-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 19-Jun-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.134 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5058524 | 22-May-2025 | Solution Available | Superseded by KB5060954 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5058383 | 22-May-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 22-May-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-May-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.133 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5055661 | 24-Apr-2025 | Solution Available | Superseded by KB5058524 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5055521 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5055170 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5053594 | 28-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5054006 | 28-Mar-2025 | Solution Available | Superseded by KB5055661 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 28-Mar-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5052006 | 03-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 03-Mar-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.132 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5046855 | 03-Mar-2025 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5050109 | 03-Feb-2025 | Solution Available | Superseded by KB5054006 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5049993 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5049614 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 03-Feb-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.131 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5048671 | 19-Dec-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5046612 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5046266 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 25-Nov-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.130 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5044293 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5044021 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 18-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.129 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 18-Oct-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5040946 | 01-Oct-2024 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5043124 | 01-Oct-2024 | Solution Available | Superseded by KB5050109 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5043051 | 01-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 01-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x65 – v5.128 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5041576 | 06-Sep-2024 | Solution Available | Superseded by KB5043124 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5041773 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5041951 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 06-Sep-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 06-Sep-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.127 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5040562 | 04-Aug-2024 | Solution Available | Superseded by KB5041576 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5040434 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5039885 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 04-Aug-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| Kb890830* | 04-Aug-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.126 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5039334 | 08-Jul-2024 | Solution Available | Superseded by KB5040562 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5039214 | 08-Jul-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 08-Jul-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x65 – v5.125 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 08-Jul-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5037763 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5037926 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 17-Jun-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.124 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5037016 | 10-May-2024 | Solution Available | Superseded by KB5039334 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5036899 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 10-May-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.123 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5036609 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 10-May-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5035855 | 22-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 22-Apr-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5035962 | 22-Apr-2024 | Solution Available | Superseded by KB5037016 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 01-Apr-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.121 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5034767 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5034614 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5034862 | 01-Apr-2024 | Solution Available | Superseded by KB5035962 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 22-Feb-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.120 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5034119 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5033910 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5029186 | 25-Jan-2024 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5033373 | 25-Jan-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 25-Jan-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Dec-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.119 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 21-Dec-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5031989 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5032197 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5032391 | 21-Dec-2023 | Solution Available | Superseded by KB5034862 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 29-Nov-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.118 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5031362 | 29-Nov-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 29-Nov-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 26-Oct-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.117 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5030504 | 26-Oct-2023 | Solution Available | Superseded by KB5032391 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5030213 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5029924 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 26-Oct-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 14-Sep-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.116 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5029242 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5028952 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 14-Sep-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5003279 | 28-Aug-2023 | Solution Available | Install Recommended Update | SQL Server 2016 Service Pack 3 | Microsoft SQL Server 2016 |
| KB890830* | 28-Aug-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.115 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 28-Aug-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5028169 | 28-Aug-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 25-Jul-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.114 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 25-Jul-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5027219 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5027123 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 12-Jun-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.113 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 12-Jun-2023 | Solution Available | Install Recommended Update | Update for Microsoft Defender Antivirus antimalware platform (Version 4.18.2304.8) | Microsoft Defender Antivirus (platform update) |
| KB5026363 | 12-Jun-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5025228 | 11-May-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5023697 | 12-Apr-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5023788 | 12-Apr-2023 | Solution Available | Superseded by KB5030504 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5022838 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5022503 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5022289 | 07-Feb-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5021235 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5020873 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5020614 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5019964 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5018411 | 14-Dec-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5017396 | 27-Oct-2022 | Solution Available | Superseded by KB5023788 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5017305 | 27-Oct-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5017095 | 14-Sep-2022 | Solution Available | Superseded by KB5017396 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5016622 | 14-Sep-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5012170 | 14-Sep-2022 | Solution Available | Install Recommended Update | Security Update for Windows |  |
| KB5015808 | 03-Aug-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5016058 | 03-Aug-2022 | Solution Available | Superseded by KB5017095 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5014702 | 25-Jul-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5014026 | 13-Jun-2022 | Solution Available | Superseded by KB5016058 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5013952 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5013625 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5012596 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5012118 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5011570 | 18-Apr-2022 | Solution Available | Superseded by KB5014026 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5011495 | 18-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5010359 | 01-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5009546 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5008877 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5008207 | 26-Jan-2022 | Solution Available | Install Recommended Update | None | N/A |
| KB5007192 | 13-Jan-2022 | Solution Available | Install Recommended Update | None | N/A |
| KB5006669 | 30-Nov-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5005698 | 10-Nov-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5005573 | 10-Nov-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5005043 | 21-Sep-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5004948 | 29-Jul-2021 | Solution Available | Install Recommended Update | Superseded by KB5004238 | N/A |
| KB5004238 | 29-Jul-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5003638 | 07-Jul-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5003197 | 02-Jun-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5001402 | 04-May-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5001347 | 04-May-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5000803 | 02-Apr-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5001078 | 03-Mar-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4601318 | 03-Mar-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4601051 | 03-Mar-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4598243 | 10-Feb-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4535680 | 10-Feb-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4593226 | 10-Feb-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4586830 | 10-Dec-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4580346 | 03-Nov-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4580325 | 03-Nov-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4577015 | 22-Oct-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4576750 | 22-Oct-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4571694 | 12-Oct-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4565912 | 29-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4565511 | 29-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4562561 | 13-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4561616 | 13-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4561600 | 13-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4556813 | 08-Jun-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4550994 | 12-May-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4550929 | 12-May-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4540670 | 15-Apr-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4540723 | 15-Apr-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4537764 | 13-Mar-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4537759 | 13-Mar-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4534271 | 06-Feb-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4530689 | 08-Jan-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4525236 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4516115 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4520724 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493478 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4519998 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4516115 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4521858 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493478 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4516115 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4516044 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4512574 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493478 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4512517 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4503308 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4509091 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4103729 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4103723 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4507460 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4503308 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4509091 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4103729 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4103723 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4503537 | 02-Jul-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4503308 | 02-Jul-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4503267 | 02-Jul-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4498947 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4494440 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4497932 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4485447 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493478 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830* | 09-Apr-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4485447 | 09-Apr-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4489907 | 09-Apr-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4489882 | 09-Apr-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830* | 13-Mar-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4485447 | 13-Mar-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4487038 | 13-Mar-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4487026 | 13-Mar-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4465659 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4480961 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4480979 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4465659 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4471321 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4471331 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4457146 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4467691 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4467694 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4465659 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4462917 | 06-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4457146 | 06-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 06-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 06-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4457131 | 28-Sep-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4457146 | 28-Sep-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 28-Sep-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 28-Sep-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4343887 | 30-Aug-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4343902 | 30-Aug-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 30-Aug-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 30-Aug-2018 | Solution Available | Install Recommended Update | None | N/A |

## Windows Server 2016

| Release: | PIC iX C.0x | OPERATING_SYSTEM / SERVICE_PACK: (The Main Operating System is the OS of the system component that connects to the network) | PIC iX Servers and Client Application: Windows Server 2016 |
| --- | --- | --- | --- |

| Windows Server 2016 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Vulnerability / Patch ID | Last Update by Philips | Activity Status | Recommended Customer Action | Notes / Instructions | Microsoft Product |
| KB5088064 | 27-May-2026 | Solution Available | Install Recommended Update prior to other updates | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5087537 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5087065 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 27-May-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.141 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077473** | 27-May-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 CU | Microsoft SQL Server 2016 |
| KB5082089 | 21-Apr-2026 | Solution Available | Superseded by KB5088064 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5082198 | **Only for customers using the Azure Connect Feature update path. 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5082411 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 21-Apr-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Apr-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.140 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077474 | 21-Apr-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5078938 | 20-Mar-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5075902 | 23-Feb-2026 | Solution Available | Superseded by KB5082089 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5075999 | 23-Feb-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 23-Feb-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 23-Feb-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.139 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5073447 | 26-Jan-2026 | Solution Available | Superseded by KB5075902 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5073722 | 26-Jan-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5068401 | 26-Jan-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5071543 | 17-Dec-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 17-Dec-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 17-Dec-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool - v5.138 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5070247 | 25-Nov-2025 | Solution Available | Superseded by KB5073447 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5068864 | 25-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 25-Nov-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool - v5.137 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5070882 | 12-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows Server 2016 |
| KB5066584 | 22-Oct-2025 | Solution Available | Superseded by KB5070247 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5066836 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5066136 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 22-Oct-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Oct-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool - v5.136 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5065687 | 17-Sep-2025 | Solution Available | Superseded by KB5066584 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5065427 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5065749 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5063871 | 20-Aug-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 20-Aug-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.135 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 20-Aug-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5058718 | 20-Aug-2025 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5062799 | 20-Jul-2025 | Solution Available | Superseded by KB5065687 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5062560 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5062064 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5060954 | 19-Jun-2025 | Solution Available | Superseded by KB5062799 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5061010 | 19-Jun-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 19-Jun-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 19-Jun-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.134 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5058524 | 22-May-2025 | Solution Available | Install Recommended Update prior to other updates | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5058383 | 22-May-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 22-May-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-May-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.133 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5055661 | 24-Apr-2025 | Solution Available | Superseded by KB5058524 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5055521 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5055170 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5053594 | 28-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5054006 | 28-Mar-2025 | Solution Available | Superseded by KB5055661 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 28-Mar-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5052006 | 03-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 03-Mar-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.132 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5046855 | 03-Mar-2025 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5050109 | 03-Feb-2025 | Solution Available | Superseded by KB5054006 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5049993 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5049614 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 03-Feb-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.131 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5048671 | 19-Dec-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5046612 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5046266 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 25-Nov-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.130 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5044293 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5044021 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 18-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.129 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 18-Oct-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5043124 | 01-Oct-2024 | Solution Available | Superseded by KB5050109 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5043051 | 01-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 01-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.128 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5040946 | 01-Oct-2024 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5041576 | 06-Sep-2024 | Solution Available | Superseded by KB5043124 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5041773 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5041951 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 06-Sep-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 06-Sep-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.127 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5040562 | 04-Aug-2024 | Solution Available | Superseded by KB5041576 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5040434 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5039885 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 04-Aug-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| Kb890830* | 04-Aug-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.126 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5039334 | 08-Jul-2024 | Solution Available | Superseded by KB5040562 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5039214 | 08-Jul-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 08-Jul-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.125 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 08-Jul-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5037763 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5037926 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 17-Jun-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.124 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5037016 | 10-May-2024 | Solution Available | Superseded by KB5039334 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5036899 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 10-May-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.123 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5036609 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 10-May-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5035855 | 22-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 22-Apr-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5035962 | 22-Apr-2024 | Solution Available | Superseded by KB5037016 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 01-Apr-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.121 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5034767 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5034614 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5034862 | 01-Apr-2024 | Solution Available | Superseded by KB5035962 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 22-Feb-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.120 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5034119 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5033910 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5029186 | 25-Jan-2024 | Solution Available | Install Recommended Update | Security Update for SQL Server 2016 Service Pack 3 GDR | Microsoft SQL Server 2016 |
| KB5033373 | 25-Jan-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 25-Jan-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Dec-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.119 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 21-Dec-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5031989 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5032197 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5032391 | 21-Dec-2023 | Solution Available | Superseded by KB5034862 | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 29-Nov-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.118 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5031362 | 29-Nov-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 29-Nov-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 26-Oct-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.117 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5030504 | 26-Oct-2023 | Solution Available | Install Recommended Update prior to other updates | Servicing Stack Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5030213 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5029924 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 26-Oct-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 14-Sep-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.116 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5029242 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5028952 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB4052623 | 14-Sep-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5003279 | 28-Aug-2023 | Solution Available | Install Recommended Update | SQL Server 2016 Service Pack 3 | Microsoft SQL Server 2016 |
| KB890830* | 28-Aug-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.115 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 28-Aug-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5028169 | 28-Aug-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 25-Jul-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.114 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 25-Jul-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5027219 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5027123 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB890830* | 12-Jun-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.113 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 12-Jun-2023 | Solution Available | Install Recommended Update | Update for Microsoft Defender Antivirus antimalware platform (Version 4.18.2304.8) | Microsoft Defender Antivirus (platform update) |
| KB5026363 | 12-Jun-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5025228 | 11-May-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5023697 | 12-Apr-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5023788 | 12-Apr-2023 | Solution Available | Install Recommended Update prior to other updates | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5022838 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5022503 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5022289 | 07-Feb-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5021235 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5020873 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5020614 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5019964 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5018411 | 14-Dec-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5017396 | 27-Oct-2022 | Solution Available | Superseded by KB5023788 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5017305 | 27-Oct-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5017095 | 14-Sep-2022 | Solution Available | Superseded by KB5017396 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5016622 | 14-Sep-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5012170 | 14-Sep-2022 | Solution Available | Install Recommended Update | Security Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5015808 | 03-Aug-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5016058 | 03-Aug-2022 | Solution Available | Superseded by KB5017095 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5014702 | 25-Jul-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5014026 | 13-Jun-2022 | Solution Available | Superseded by KB5016058 | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5013952 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5013625 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5012596 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5012118 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5011570 | 18-Apr-2022 | Solution Available | Install Recommended Update prior to other updates | Servicing Stack Update | Windows 10 Version 1607 Windows Server 2016 |
| KB5011495 | 18-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5010359 | 01-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5009546 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1607 Windows Server 2016 |
| KB5008877 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1607 Windows Server 2016 |
| KB5008207 | 26-Jan-2022 | Solution Available | Install Recommended Update | None | N/A |
| KB5007192 | 13-Jan-2022 | Solution Available | Install Recommended Update | None | N/A |
| KB5006669 | 30-Nov-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5005698 | 10-Nov-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5005573 | 10-Nov-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5005043 | 21-Sep-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5004948 | 29-Jul-2021 | Solution Available | Install Recommended Update | Superseded by KB5004238 | N/A |
| KB5004238 | 29-Jul-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5003638 | 07-Jul-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5003197 | 02-Jun-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5001402 | 04-May-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5001347 | 04-May-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5000803 | 02-Apr-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB5001078 | 03-Mar-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4601318 | 03-Mar-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4601051 | 03-Mar-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4598243 | 10-Feb-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4535680 | 10-Feb-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4593226 | 10-Feb-2021 | Solution Available | Install Recommended Update | None | N/A |
| KB4586830 | 10-Dec-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4580346 | 03-Nov-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4577015 | 22-Oct-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4576750 | 22-Oct-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4571694 | 12-Oct-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4565912 | 29-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4565511 | 29-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4562561 | 13-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4561616 | 13-Jul-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4556813 | 08-Jun-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4550994 | 12-May-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4550929 | 12-May-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4540670 | 15-Apr-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4540723 | 15-Apr-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4537764 | 13-Mar-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4534271 | 06-Feb-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4530689 | 08-Jan-2020 | Solution Available | Install Recommended Update | None | N/A |
| KB4525236 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4520724 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 05-Dec-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4519998 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4521858 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 29-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4516044 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4512574 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 03-Oct-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4512517 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4509091 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4103723 | 26-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4507460 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4509091 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4103723 | 02-Aug-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4503537 | 02-Jul-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4503267 | 02-Jul-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4498947 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4494440 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4485447 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4493470 | 21-May-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830* | 09-Apr-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4485447 | 09-Apr-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4489882 | 09-Apr-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830* | 13-Mar-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4485447 | 13-Mar-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4487026 | 13-Mar-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4465659 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4480961 | 31-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4465659 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4471321 | 02-Jan-2019 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4467691 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4465659 | 30-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4462917 | 06-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 06-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 06-Nov-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4457131 | 28-Sep-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 28-Sep-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 28-Sep-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4343887 | 30-Aug-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB890830 | 30-Aug-2018 | Solution Available | Install Recommended Update | None | N/A |
| KB4132216 | 30-Aug-2018 | Solution Available | Install Recommended Update | None | N/A |

## Windows 10 Enterprise LTSC v1809

| Release: | PIC iX 4.x | OPERATING_SYSTEM / SERVICE_PACK: (The Main Operating System is the OS of the system component that connects to the network) | PIC iX Servers and Client Application: Windows 10 Enterprise LTSC v1809 |
| --- | --- | --- | --- |

| Windows 10 Enterprise LTSC v1809 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Vulnerability / Patch ID | Last Update by Philips | Activity Status | Recommended Customer Action | Notes / Instructions | Microsoft Product |
| KB5087538 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5088864 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 27-May-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.141 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5093447 | 27-May-2026 | Solution Available | Install: dotnet-hosting-8.0.27-win.exe for x64 Client windowsdesktop-runtime-8.0.27-win-x64.exe | .NET 8.0.27 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 27-May-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 148 Update for x64 based Editions (Build 148.0.3967.54) Microsoft Edge-WebView2 Runtime Version 148 Update for x64 based Editions (Build 148.0.3967.54) | Microsoft Edge (Build 148.0.3967.54) | Microsoft Edge |
| KB5082123 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5084066 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 21-Apr-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Apr-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.140 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5086096 | 21-Apr-2026 | Solution Available | Install: dotnet-hosting-8.0.26-win.exe for x64 Client windowsdesktop-runtime-8.0.26-win-x64.exe | .NET 8.0.26 Update for x86 Client | .NET 8.0 |
| Microsoft Edge | 21-Apr-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 147 Update for x64 based Editions (Build 147.0.3912.60) Microsoft Edge-WebView2 Runtime Version 145 Update for x64 based Editions (Build 147.0.3912.60) | Microsoft Edge (Build 147.0.391.60) | Microsoft Edge |
| KB5077469 | 21-Apr-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2019 RTM CU | Microsoft SQL Server 2019 |
| KB5078752 | 20-Mar-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5081277 | 20-Mar-2026 | Solution Available | Install: dotnet-hosting-8.0.25-win.exe for x64 Client windowsdesktop-runtime-8.0.25-win-x64.exe | .NET 8.0.25 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 20-Mar-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 145 Update for x64 based Editions (Build 145.0.3800.97) Microsoft Edge-WebView2 Runtime Version 145 Update for x64 based Editions (Build 145.3800.97) | Microsoft Edge (Build 145.0.3800.97) | Microsoft Edge |
| 7-Zip | 20-Mar-2026 | Solution Available | Install Recommended Update | 7-Zip Version 26.00 | 7-Zip |
| CVE-2025-29803 | 20-Mar-2026 | Solution Available | Install: Visual Studio Tools for Applications (VSTA) 2017: Version 15.0.36010 Visual Studio Tools for Applications (VSTA) 2019: Version 16.0.35907 | Visual Studio Tools for Applications and SQL Server Management Studio Elevation of Privilege Vulnerability | Windows 10 Version 1809 Windows Server 2019 |
| KB5075904 | 23-Feb-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 23-Feb-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 23-Feb-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.139 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077863 | 23-Feb-2026 | Solution Available | Install: dotnet-hosting-8.0.24-win.exe windowsdesktop-runtime-8.0.24-win-x64.exe | .NET 8.0.24 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 23-Feb-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 144 Update for x64 based Editions (Build 144.0.3719.115) Microsoft Edge-WebView2 Runtime Version 144 Update for x64 based Editions (Build 144.0.3719.115) | Microsoft Edge (Build 144.0.3719.115) | Microsoft Edge |
| KB5068404 | 23-Feb-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2019 RTM CU | Microsoft SQL Server 2019 |
| KB5073723 | 26-Jan-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5074752 | 26-Jan-2026 | Solution Available | Install: dotnet-hosting-8.0.23-win.exe windowsdesktop-runtime-8.0.23-win-x64.exe | .NET 8.0.23 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 26-Jan-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.139) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build 143.0.3650.139) | Microsoft Edge (Build 143.0.3650.139) | Microsoft Edge |
| KB5071544 | 17-Dec-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 17-Dec-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 17-Dec-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.138 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| Microsoft Edge | 17-Dec-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.66) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build 143.0.3650.66) | Microsoft Edge (Build 143.0.3650.66) | Microsoft Edge |
| KB5068791 | 25-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 25-Nov-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.137 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5071066 | 25-Nov-2025 | Solution Available | Install: dotnet-hosting-8.0.22-win.exe windowsdesktop-runtime-8.0.22-win-x64.exe | .NET 8.0.22 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 25-Nov-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 142 Update for x64 based Editions (Build 142.0.3595.69) Microsoft Edge-WebView2 Runtime Version 142 Update for x64 based Editions (Build 142.0.3595.69) | Microsoft Edge (Build 142.0.3595.69) | Microsoft Edge |
| KB5066586 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5066738 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 22-Oct-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Oct-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.136 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5068331 | 22-Oct-2025 | Solution Available | Install: dotnet-hosting-8.0.21-win.exe windowsdesktop-runtime-8.0.21-win-x64.exe | .NET 8.0.21 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 22-Oct-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 141 Update for x64 based Editions (Build 141.0.3537.71) Microsoft Edge-WebView2 Runtime Version 141 Update for x64 based Editions (Build 141.0.3537.71) | Microsoft Edge (Build 141.0.3537.71) | Microsoft Edge |
| 7-Zip | 22-Oct-2025 | Solution Available | Install Recommended Update | 7-Zip (Build 25.01) | 7-Zip |
| KB5065428 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5065955 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET | Windows 10 Version 1809 Windows Server 2019 |
| KB5066233 | 17-Sep-2025 | Solution Available | Install: dotnet-hosting-8.0.20-win.exe windowsdesktop-runtime-8.0.20-win-x64.exe | .NET 8.0.20 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 17-Sep-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 140 Update for x64 based Editions (Build 140.0.3485.54) Microsoft Edge-WebView2 Runtime Version 140 Update for x64 based Editions (Build 140.0.3485.54) | Microsoft Edge (Build 140.0.3485.54) | Microsoft Edge |
| KB5063877 | 20-Aug-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 20-Aug-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.1​35 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 20-Aug-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5064838 | 20-Aug-2025 | Solution Available | Install: dotnet-hosting-8.0.19-win.exe windowsdesktop-runtime-8.0.19-win-x64.exe | .NET 8.0.19 Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 20-Aug-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 139 Update for x64 based Editions (Build 139.0.3405.86) Microsoft Edge-WebView2 Runtime Version 139 Update for x64 based Editions (Build 139.0.3405.86) | Microsoft Edge (Build 139.0.3405.86) | Microsoft Edge |
| KB5062557 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5062152 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5063326 | 20-Jul-2025 | Solution Available | Install: dotnet-hosting-8.0.18-win.exe windowsdesktop-runtime-8.0.18-win-x64.exe | .NET 8.0.18 Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 20-Jul-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 138 Update for x64 based Editions (Build 138.0.3351.77) Microsoft Edge-WebView2 Runtime Version 138 Update for x64 based Editions (Build 138.0.3351.77) | Microsoft Edge (Build 138.0.3351.77) | Microsoft Edge |
| KB5060531 | 19-Jun-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 19-Jun-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 19-Jun-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.134 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5061935 | 19-Jun-2025 | Solution Available | Install: dotnet-hosting-8.0.17-win.exe windowsdesktop-runtime-8.0.17-win-x64.exe | .NET 8.0.17 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 19-Jun-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 137 Update for x64 based Editions (Build 137.0.3296.68) Microsoft Edge-WebView2 Runtime Version 137 Update for x64 based Editions (Build 137.0.3296.68) | Microsoft Edge (Build 137.0.3296.68) | Microsoft Edge |
| KB5054833 | 19-Jun-2025 | Solution Available | Install Recommended Update | SQL Server 2019 Cumulative Update (CU32) | Microsoft SQL Server 2019 |
| KB5058392 | 22-May-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 22-May-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-May-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.133 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5059200 | 22-May-2025 | Solution Available | Install: dotnet-hosting-8.0.16-win.exe windowsdesktop-runtime-8.0.16-win-x64.exe | .NET 8.0.16 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 22-May-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 136 Update for x64 based Editions (Build 136.0.3240.64) Microsoft Edge-WebView2 Runtime Version 136 Update for x64 based Editions (Build 136.0.3240.64) | Microsoft Edge (Build 136.0.3240.64) | Microsoft Edge |
| KB5055519 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5055681 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5056686 | 24-Apr-2025 | Solution Available | Install: dotnet-hosting-8.0.15-win.exe windowsdesktop-runtime-8.0.15-win-x64.exe | .NET 8.0.15 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 24-Apr-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 135 Update for x64 based Editions (Build 135.0.3179.54) Microsoft Edge-WebView2 Runtime Version 135 Update for x64 based Editions (Build 135.0.3179.54) | Microsoft Edge (Build 135.0.3179.54) | Microsoft Edge |
| KB5053596 | 28-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 28-Mar-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5054229 | 28-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.14-win.exe windowsdesktop-runtime-8.0.14-win-x64.exe | .NET 8.0.14 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 28-Mar-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 134 Update for x64 based Editions (Build 134.0.3124.51) Microsoft Edge-WebView2 Runtime Version 134 Update for x64 based Editions (Build 134.0.3124.51) | Microsoft Edge (Build 134.0.3124.51) | Microsoft Edge |
| KB5052000 | 03-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 03-Mar-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.132 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5052978 | 03-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.13-win.exe windowsdesktop-runtime-8.0.13-win-x64.exe | .NET 8.0.13 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 03-Mar-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 133 Update for x64 based Editions (Build 133.0.3065.59) Microsoft Edge-WebView2 Runtime Version 133 Update for x64 based Editions (Build 133.0.3065.59) | Microsoft Edge (Build 133.0.3065.59) | Microsoft Edge |
| KB5050008 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5050182 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 03-Feb-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.131 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5050525 | 03-Feb-2025 | Solution Available | Install: dotnet-hosting-8.0.12-win.exe windowsdesktop-runtime-8.0.12-win-x64.exe | .NET 8.0.12 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 03-Feb-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 132 Update for x64 based Editions (Build 132.0.2957.115) Microsoft Edge-WebView2 Runtime Version 132 Update for x64 based Editions (Build 132.0.2957.115) | Microsoft Edge (Build 132.0.2957.115) | Microsoft Edge |
| KB5049235 | 03-Feb-2025 | Solution Available | Install Recommended Update | SQL Server 2019 Cumulative Update (CU30) | Microsoft SQL Server 2019 |
| KB5048661 | 19-Dec-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| 7-Zip | 19-Dec-2024 | Solution Available | Install Recommended Update | 7-Zip (Build 24.09) | 7-Zip |
| KB5046615 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5046540 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 25-Nov-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.130 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5047486 | 25-Nov-2024 | Solution Available | Install: dotnet-hosting-6.0.36-win.exe windowsdesktop-runtime-6.0.36-win-x64.exe | .NET 6.0.36 Security Update for x64 Client (Applicable for PIC iX 4.0/4.1/4.2) | .NET 6.0 |
| KB5047489 | 25-Nov-2024 | Solution Available | Install: dotnet-hosting-8.0.11-win.exe windowsdesktop-runtime-8.0.11-win-x64.exe | .NET 8.0.11 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 25-Nov-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 131 Update for x64 based Editions (Build 131.0.2903.48) Microsoft Edge-WebView2 Runtime Version 131 Update for x64 based Editions (Build 131.0.2903.48) | Microsoft Edge (Build 131.0.2903.48) | Microsoft Edge |
| KB5044277 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5044089 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 18-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.129 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 18-Oct-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5045998 | 18-Oct-2024 | Solution Available | Install: dotnet-hosting-6.0.35-win-x64.exe windowsdesktop-runtime-6.0.35-win-x64.exe | .NET 6.0.35 Security Update for x64 Client (Applicable for PIC iX 4.0/4.1/4.2) | .NET 6.0 |
| KB5045993 | 18-Oct-2024 | Solution Available | Install: dotnet-hosting-8.0.10-win-x64.exe windowsdesktop-runtime-8.0.10-win-x64.exe | .NET 8.0.10 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 18-Oct-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 129 Update for x64 based Editions (Build 129.0.2792.52) Microsoft Edge-WebView2 Runtime Version 129 Update for x64 based Editions | Microsoft Edge (Build 129.0.2792.52) | Microsoft Edge |
| KB5043050 | 01-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 01-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.128 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5042132 | 01-Oct-2024 | Solution Available | Install: dotnet-hosting-8.0.8-win.exe windowsdesktop-runtime-8.0.8-win-x64.exe | .NET 8.0.8 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 01-Oct-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 128 Update for x64 based Editions (Build 128.0.2739.42) Microsoft Edge-WebView2 Runtime Version 128 Update for x64 based Editions (Build 128.0.2739.42) | Microsoft Edge (Build 128.0.2739.42) | Microsoft Edge |
| KB5041578 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5042350 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 06-Sep-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 06-Sep-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.127 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5042131 | 06-Sep-2024 | Solution Available | Install: dotnet-hosting-6.0.33-win.exe windowsdesktop-runtime-6.0.33-win-x64.exe | .NET 6.0.33 Security Update for x64 Client | .NET 6.0 |
| Microsoft Edge | 06-Sep-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 127 Update for x64 based Editions (Build 127.0.2651.74) Microsoft Edge-WebView2 Runtime Version 127 Update for x64 based Editions (Build 127.0.2651.74) | Microsoft Edge (Build 127.0.2651.74) | Microsoft Edge |
| KB5039747 | 06-Sep-2024 | Solution Available | Install Recommended Update | SQL Server 2019 Cumulative Update (CU) 28 | Microsoft SQL Server 2019 |
| KB5040430 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5041017 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 04-Aug-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5041080 | 04-Aug-2024 | Solution Available | Install: dotnet-hosting-6.0.32-win.exe windowsdesktop-runtime-6.0.32-win-x64.exe | .NET 6.0.32 Security Update for x64 Client | .NET 6.0 |
| KB890830* | 04-Aug-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.126 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5039217 | 08-Jul-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5039843 | 08-Jul-2024 | Solution Available | Install: dotnet-hosting-6.0.31-win.exe windowsdesktop-runtime-6.0.31-win-x64.exe | .NET 6.0.31 Security Update for x64 Client | .NET 6.0 |
| KB890830* | 08-Jul-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.125 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 08-Jul-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| Microsoft Edge | 08-Jul-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 126 Update for x64 based Editions (Build 126.0.2592.56) Microsoft Edge-WebView2 Runtime Version 126 Update for x64 based Editions (Build 126.0.2592.56) | Microsoft Edge (Build 126.0.2592.56) | Microsoft Edge |
| KB5037765 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5038350 | 17-Jun-2024 | Solution Available | Install: dotnet-hosting-6.0.30-win.exe windowsdesktop-runtime-6.0.30-win-x64.exe | .NET 6.0.30 Security Update for x64 Client | .NET 6.0 |
| KB5038283 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 17-Jun-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – x5.124 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| Microsoft Edge | 17-Jun-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 125 Update for x64 based Editions (Build 125.0.2535.51) Microsoft Edge-WebView2 Runtime Version 125 Update for x64 based Editions (Build 125.0.2535.51) | Microsoft Edge (Build 125.0.2535.51) | Microsoft Edge |
| KB5036896 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5037336 | 10-May-2024 | Solution Available | Install: dotnet-hosting-6.0.29-win-x64.exe windowsdesktop-runtime-6.0.29-win-x64.exe | .NET 6.0.29 Security Update for x64 Client | .NET 6.0 |
| KB890830* | 10-May-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.123 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5037034 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 10-May-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| Microsoft Edge | 10-May-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 123 Update for x64 based Editions (Build 123.0.2420.65) Microsoft Edge-WebView2 Runtime Version 123 Update for x64 based Editions (Build 123.0.2420.65) | Microsoft Edge (Build 123.0.2420.65) | Microsoft Edge |
| Microsoft Edge | 22-Apr-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 123 Update for x64 based Editions (Build 123.0.2420.65) Microsoft Edge-WebView2 Runtime Version 123 Update for x64 based Editions (Build 123.0.2420.65) | Microsoft Edge (Build 123.0.2420.65) | Microsoft Edge |
| Microsoft Edge | 22-Apr-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 122 Update for x64 based Editions (Build 122.0.2365.52) Microsoft Edge-WebView2 Runtime Version 122 Update for x64 based Editions (Build 122.0.2365.52) | Microsoft Edge (Build 122.0.2365.52) | Microsoft Edge |
| KB4052623 | 22-Apr-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Apr-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.122 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5036450 | 22-Apr-2024 | Solution Available | Install: dotnet-hosting-6.0.28-win.exe windowsdesktop-runtime-6.0.28-win-x64.exe | .NET 6.0.28 Security Update for x64 Client | .NET 6.0 |
| KB5035849 | 22-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| 7-Zip | 01-Apr-2024 | Solution Available | Install Recommended Update | 7-Zip (Build 23.01) | 7-Zip |
| KB5035119 | 01-Apr-2024 | Solution Available | Install: dotnet-hosting-6.0.27-win.exe windowsdesktop-runtime-6.0.27-win-x64.exe | .NET 6.0.27 Security Update for x64 Client | .NET 6.0 |
| KB890830* | 01-Apr-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.121 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5034683 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5034768 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5031908 | 01-Apr-2024 | Solution Available | Install Recommended Update | SQL Server 2019 RTM Cumulative Update (CU) 24 | Microsoft SQL Server 2019 |
| Microsoft Edge | 22-Feb-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 121 Update for x64 based Editions (Build 121.0.2277.83) Microsoft Edge-WebView2 Runtime Version 121 Update for x64 based Editions (Build 121.0.2277.83) | Microsoft Edge (Build 121.0.2277.83) | Microsoft Edge |
| KB890830* | 22-Feb-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.120 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5033733 | 22-Feb-2024 | Solution Available | Install: dotnet-hosting-6.0.26-win.exe windowsdesktop-runtime-6.0.26-win-x64.exe | .NET 6.0.26 Security Update for x64 Client | .NET 6.0 |
| KB5034273 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5034127 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5030333 | 25-Jan-2024 | Solution Available | Install Recommended Update | SQL Server 2019 RTM Cumulative Update (CU) 23 | Microsoft SQL Server 2019 |
| KB5033371 | 25-Jan-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 25-Jan-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Dec-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.119 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 21-Dec-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5032337 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5032196 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5032883 | 21-Dec-2023 | Solution Available | Install: dotnet-hosting-6.0.25-win.exe windowsdesktop-runtime-6.0.25-win-x64.exe | .NET 6.0.25 Security Update for x64 Client | .NET 6.0 |
| KB890830* | 29-Nov-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.118 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5032874 | 29-Nov-23 | Solution Available | Install: dotnet-hosting-6.0.24-win.exe windowsdesktop-runtime-6.0.24-win-x64.exe | .NET 6.0.24 Security Update for x64 Client | .NET 6.0 |
| KB5031361 | 29-Nov-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 29-Nov-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 26-Oct-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.117 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5030559 | 26-Oct-2023 | Solution Available | Install: dotnet-hosting-6.0.22-win windowsdesktop-runtime-6.0.22-win-x64 | .NET 6.0.22 Security Update for x64 Client | .NET 6.0 |
| KB5021953 | 26-Oct-2023 | Solution Available | Install: dotnet-hosting-3.1.32-win | .NET Core 3.1.32 Security Update for x64 Client | .NET Core 3.1 |
| KB5030214 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5030178 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 26-Oct-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 14-Sep-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.116 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5029247 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5029647 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 14-Sep-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5024276 | 28-Aug-2023 | Solution Available | Install Recommended Update | SQL Server 2019 RTM Cumulative Update (CU) 20 | Microsoft SQL Server 2019 |
| KB890830* | 28-Aug-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.115 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 28-Aug-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5028168 | 28-Aug-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 25-Jul-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.114 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 25-Jul-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5027536 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5027222 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 12-Jun-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.113 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 12-Jun-2023 | Solution Available | Install Recommended Update | Update for Microsoft Defender Antivirus antimalware platform (Version 4.18.2304.8) | Microsoft Defender Antivirus (platform update) |
| KB5026362 | 12-Jun-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5025229 | 11-May-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5023702 | 12-Apr-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5022840 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5022782 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5022286 | 07-Feb-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5021237 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5021085 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5020685 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5019966 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5018419 | 14-Dec-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5017315 | 27-Oct-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5016623 | 14-Sep-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5012170 | 14-Sep-2022 | Solution Available | Install Recommended Update | Security Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5015811 | 03-Aug-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5014692 | 25-Jul-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5013941 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5013868 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5012647 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5012328 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5011503 | 18-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5010351 | 01-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5009718 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5009557 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5005112 | 25-Feb-2022 | Solution Available | Install Recommended Update prior to other updates | Service Stack Update | Windows 10 Version 1809 Windows Server 2019 |

## Windows Server 2019

| Release: | PIC iX 4.x | OPERATING_SYSTEM / SERVICE_PACK: (The Main Operating System is the OS of the system component that connects to the network) | PIC iX Servers and Client Application: Windows Server 2019 |
| --- | --- | --- | --- |

| Windows Server 2019 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Vulnerability / Patch ID | Last Update by Philips | Activity Status | Recommended Customer Action | Notes / Instructions | Microsoft Product |
| KB5087538 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5088864 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 27-May-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.141 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5093447 | 27-May-2026 | Solution Available | Install: dotnet-hosting-8.0.27-win.exe windowsdesktop-runtime-8.0.27-win-x64.exe | .NET 8.0.27 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 27-May-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 148 Update for x64 based Editions (Build 148.0.3967.54) Microsoft Edge-WebView2 Runtime Version 148 Update for x64 based Editions (Build 148.0.3967.54) | Microsoft Edge (Build 148.0.3967.54) | Microsoft Edge |
| KB5082123 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5084066 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 21-Apr-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Apr-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.140 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5086096 | 21-Apr-2026 | Solution Available | Install: dotnet-hosting-8.0.26-win.exe for x64 Server windowsdesktop-runtime-8.0.26-win-x64.exe | .NET 8.0.26 Update for x86 Server | .NET 8.0 |
| Microsoft Edge | 21-Apr-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 147 Update for x64 based Editions (Build 147.0.3912.60) Microsoft Edge-WebView2 Runtime Version 145 Update for x64 based Editions (Build 147.0.3912.60) | Microsoft Edge (Build 147.0.391.60) | Microsoft Edge |
| KB5077469 | 21-Apr-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2019 RTM CU | Microsoft SQL Server 2019 |
| KB5078752 | 20-Mar-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5081277 | 20-Mar-2026 | Solution Available | Install: dotnet-hosting-8.0.25-win.exe for x64 Server windowsdesktop-runtime-8.0.25-win-x64.exe | .NET 8.0.25 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 20-Mar-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 145 Update for x64 based Editions (Build 145.0.3800.97) Microsoft Edge-WebView2 Runtime Version 145 Update for x64 based Editions (Build 145.3800.97) | Microsoft Edge (Build 145.0.3800.97) | Microsoft Edge |
| 7-Zip | 20-Mar-2026 | Solution Available | Install Recommended Update | 7-Zip Version 26.00 | 7-Zip |
| CVE-2025-29803 | 20-Mar-2026 | Solution Available | Install: Visual Studio Tools for Applications (VSTA) 2017: Version 15.0.36010 Visual Studio Tools for Applications (VSTA) 2019: Version 16.0.35907 | Visual Studio Tools for Applications and SQL Server Management Studio and SQL Server Elevation of Privilege Vulnerability | Windows 10 Version 1809 Windows Server 2019 |
| KB5075904 | 23-Feb-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 23-Feb-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 23-Feb-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.139 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077863 | 23-Feb-2026 | Solution Available | Install: dotnet-hosting-8.0.24-win.exe windowsdesktop-runtime-8.0.24-win-x64.exe | .NET 8.0.24 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 23-Feb-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 144 Update for x64 based Editions (Build 144.0.3719.115) Microsoft Edge-WebView2 Runtime Version 144 Update for x64 based Editions (Build 144.0.3719.115) | Microsoft Edge (Build 144.0.3719.115) | Microsoft Edge |
| KB5068404 | 23-Feb-2026 | Solution Available | Install Recommended Update | Security Update for SQL Server 2019 RTM CU | Microsoft SQL Server 2019 |
| KB5073723 | 26-Jan-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5074752 | 26-Jan-2026 | Solution Available | Install: dotnet-hosting-8.0.23-win.exe windowsdesktop-runtime-8.0.23-win-x64.exe | .NET 8.0.23 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 26-Jan-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.139) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build 143.0.3650.139) | Microsoft Edge (Build 143.0.3650.139) | Microsoft Edge |
| KB5071544 | 17-Dec-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 17-Dec-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 17-Dec-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.138 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| Microsoft Edge | 17-Dec-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.66) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build 143.0.3650.66) | Microsoft Edge (Build 143.0.3650.66) | Microsoft Edge |
| KB5068791 | 25-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 25-Nov-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.137 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5071066 | 25-Nov-2025 | Solution Available | Install: dotnet-hosting-8.0.22-win.exe windowsdesktop-runtime-8.0.22-win-x64.exe | .NET 8.0.22 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 25-Nov-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 142 Update for x64 based Editions (Build 142.0.3595.69) Microsoft Edge-WebView2 Runtime Version 142 Update for x64 based Editions (Build 142.0.3595.69) | Microsoft Edge (Build 142.0.3595.69) | Microsoft Edge |
| KB5070883 | 12-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows Server 2019 |
| KB5066586 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5066738 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 22-Oct-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Oct-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.136 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5068331 | 22-Oct-2025 | Solution Available | Install: dotnet-hosting-8.0.21-win.exe windowsdesktop-runtime-8.0.21-win-x64.exe | .NET 8.0.21 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 22-Oct-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 141 Update for x64 based Editions (Build 141.0.3537.71) Microsoft Edge-WebView2 Runtime Version 141 Update for x64 based Editions (Build 141.0.3537.71) | Microsoft Edge (Build 141.0.3537.71) | Microsoft Edge |
| 7-Zip | 22-Oct-2025 | Solution Available | Install Recommended Update | 7-Zip (Build 25.01) | 7-Zip |
| KB5065428 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5065955 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET | Windows 10 Version 1809 Windows Server 2019 |
| KB5066233 | 17-Sep-2025 | Solution Available | Install: dotnet-hosting-8.0.20-win.exe windowsdesktop-runtime-8.0.20-win-x64.exe | .NET 8.0.20 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 17-Sep-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 140 Update for x64 based Editions (Build 140.0.3485.54) Microsoft Edge-WebView2 Runtime Version 140 Update for x64 based Editions (Build 140.0.3485.54) | Microsoft Edge (Build 140.0.3485.54) | Microsoft Edge |
| KB5063877 | 20-Aug-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 20-Aug-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.1​35 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 20-Aug-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5064838 | 20-Aug-2025 | Solution Available | Install: dotnet-hosting-8.0.19-win.exe windowsdesktop-runtime-8.0.19-win-x64.exe | .NET 8.0.19 Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 20-Aug-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 139 Update for x64 based Editions (Build 139.0.3405.86) Microsoft Edge-WebView2 Runtime Version 139 Update for x64 based Editions (Build 139.0.3405.86) | Microsoft Edge (Build 139.0.3405.86) | Microsoft Edge |
| KB5062557 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5062152 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5063326 | 20-Jul-2025 | Solution Available | Install: dotnet-hosting-8.0.18-win.exe windowsdesktop-runtime-8.0.18-win-x64.exe | .NET 8.0.18 Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 20-Jul-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 138 Update for x64 based Editions (Build 138.0.3351.77) Microsoft Edge-WebView2 Runtime Version 138 Update for x64 based Editions (Build 138.0.3351.77) | Microsoft Edge (Build 138.0.3351.77) | Microsoft Edge |
| KB5060531 | 19-Jun-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 19-Jun-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 19-Jun-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.134 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5061935 | 19-Jun-2025 | Solution Available | Install: dotnet-hosting-8.0.17-win.exe windowsdesktop-runtime-8.0.17-win-x64.exe | .NET 8.0.17 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 19-Jun-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 137 Update for x64 based Editions (Build 137.0.3296.68) Microsoft Edge-WebView2 Runtime Version 137 Update for x64 based Editions (Build 137.0.3296.68) | Microsoft Edge (Build 137.0.3296.68) | Microsoft Edge |
| KB5054833 | 19-Jun-2025 | Solution Available | Install Recommended Update | SQL Server 2019 Cumulative Update (CU32) | Microsoft SQL Server 2019 |
| KB5058392 | 22-May-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 22-May-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-May-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.133 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5059200 | 22-May-2025 | Solution Available | Install: dotnet-hosting-8.0.16-win.exe windowsdesktop-runtime-8.0.16-win-x64.exe | .NET 8.0.16 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 22-May-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 136 Update for x64 based Editions (Build 136.0.3240.64) Microsoft Edge-WebView2 Runtime Version 136 Update for x64 based Editions (Build 136.0.3240.64) | Microsoft Edge (Build 136.0.3240.64) | Microsoft Edge |
| KB5055519 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5055681 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5056686 | 24-Apr-2025 | Solution Available | Install: dotnet-hosting-8.0.15-win.exe windowsdesktop-runtime-8.0.15-win-x64.exe | .NET 8.0.15 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 24-Apr-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 135 Update for x64 based Editions (Build 135.0.3179.54) Microsoft Edge-WebView2 Runtime Version 135 Update for x64 based Editions (Build 135.0.3179.54) | Microsoft Edge (Build 135.0.3179.54) | Microsoft Edge |
| KB5053596 | 28-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 28-Mar-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5054229 | 28-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.14-win.exe windowsdesktop-runtime-8.0.14-win-x64.exe | .NET 8.0.14 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 28-Mar-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 134 Update for x64 based Editions (Build 134.0.3124.51) Microsoft Edge-WebView2 Runtime Version 134 Update for x64 based Editions (Build 134.0.3124.51) | Microsoft Edge (Build 134.0.3124.51) | Microsoft Edge |
| KB5052000 | 03-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 03-Mar-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.132 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5052978 | 03-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.13-win.exe windowsdesktop-runtime-8.0.13-win-x64.exe | .NET 8.0.13 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 03-Mar-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 133 Update for x64 based Editions (Build 133.0.3065.59) Microsoft Edge-WebView2 Runtime Version 133 Update for x64 based Editions (Build 133.0.3065.59) | Microsoft Edge (Build 133.0.3065.59) | Microsoft Edge |
| KB5050008 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5050182 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 03-Feb-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.131 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5050525 | 03-Feb-2025 | Solution Available | Install: dotnet-hosting-8.0.12-win.exe windowsdesktop-runtime-8.0.12-win-x64.exe | .NET 8.0.12 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 03-Feb-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 132 Update for x64 based Editions (Build 132.0.2957.115) Microsoft Edge-WebView2 Runtime Version 132 Update for x64 based Editions (Build 132.0.2957.115) | Microsoft Edge (Build 132.0.2957.115) | Microsoft Edge |
| KB5049235 | 03-Feb-2025 | Solution Available | Install Recommended Update | SQL Server 2019 Cumulative Update (CU30) | Microsoft SQL Server 2019 |
| KB5048661 | 19-Dec-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| 7-Zip | 19-Dec-2024 | Solution Available | Install Recommended Update | 7-Zip (Build 24.09) | 7-Zip |
| KB5046615 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5046540 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 25-Nov-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.130 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5047486 | 25-Nov-2024 | Solution Available | Install: dotnet-hosting-6.0.36-win.exe windowsdesktop-runtime-6.0.36-win-x64.exe | .NET 6.0.36 Security Update for x64 Server (Applicable for PIC iX 4.0/4.1/4.2) | .NET 6.0 |
| KB5047489 | 25-Nov-2024 | Solution Available | Install: dotnet-hosting-8.0.11-win.exe windowsdesktop-runtime-8.0.11-win-x64.exe | .NET 8.0.11 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 25-Nov-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 131 Update for x64 based Editions (Build 131.0.2903.48) Microsoft Edge-WebView2 Runtime Version 131 Update for x64 based Editions (Build 131.0.2903.48) | Microsoft Edge (Build 131.0.2903.48) | Microsoft Edge |
| KB5044277 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5044089 | 18-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 18-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.129 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 18-Oct-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5045998 | 18-Oct-2024 | Solution Available | Install: dotnet-hosting-6.0.35-win-x64.exe windowsdesktop-runtime-6.0.35-win-x64.exe | .NET 6.0.35 Security Update for x64 Server | .NET 6.0 |
| KB5045993 | 18-Oct-2024 | Solution Available | Install: dotnet-hosting-8.0.10-win-x64.exe windowsdesktop-runtime-8.0.10-win-x64.exe | .NET 8.0.10 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 18-Oct-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 129 Update for x64 based Editions (Build 129.0.2792.52) Microsoft Edge-WebView2 Runtime Version 129 Update for x64 based Editions | Microsoft Edge (Build 129.0.2792.52) | Microsoft Edge |
| KB5043050 | 01-Oct-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 01-Oct-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.128 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5042132 | 01-Oct-2024 | Solution Available | Install: dotnet-hosting-8.0.8-win.exe windowsdesktop-runtime-8.0.8-win-x64.exe | .NET 8.0.8 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 01-Oct-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 128 Update for x64 based Editions (Build 128.0.2739.42) Microsoft Edge-WebView2 Runtime Version 128 Update for x64 based Editions (Build 128.0.2739.42) | Microsoft Edge (Build 128.0.2739.42) | Microsoft Edge |
| KB5041578 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5042350 | 06-Sep-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 06-Sep-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 06-Sep-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.127 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5042131 | 06-Sep-2024 | Solution Available | Install: dotnet-hosting-6.0.33-win.exe windowsdesktop-runtime-6.0.33-win-x64.exe | .NET 6.0.33 Security Update for x64 Server | .NET 6.0 |
| Microsoft Edge | 06-Sep-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 127 Update for x64 based Editions (Build 127.0.2651.74) Microsoft Edge-WebView2 Runtime Version 127 Update for x64 based Editions (Build 127.0.2651.74) | Microsoft Edge (Build 127.0.2651.74) | Microsoft Edge |
| KB5039747 | 06-Sep-2024 | Solution Available | Install Recommended Update | SQL Server 2019 Cumulative Update (CU) 28 | Microsoft SQL Server 2019 |
| KB5040430 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5041017 | 04-Aug-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 04-Aug-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5041080 | 04-Aug-2024 | Solution Available | Install: dotnet-hosting-6.0.32-win.exe windowsdesktop-runtime-6.0.32-win-x64.exe | .NET 6.0.32 Security Update for x64 Server | .NET 6.0 |
| KB890830* | 04-Aug-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.126 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5039217 | 08-Jul-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5039843 | 08-Jul-2024 | Solution Available | Install: dotnet-hosting-6.0.31-win.exe windowsdesktop-runtime-6.0.31-win-x64.exe | .NET 6.0.31 Security Update for x64 Server | .NET 6.0 |
| KB890830* | 08-Jul-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.125 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 08-Jul-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| Microsoft Edge | 08-Jul-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 126 Update for x64 based Editions (Build 126.0.2592.56) Microsoft Edge-WebView2 Runtime Version 126 Update for x64 based Editions (Build 126.0.2592.56) | Microsoft Edge (Build 126.0.2592.56) | Microsoft Edge |
| KB5037765 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5038350 | 17-Jun-2024 | Solution Available | Install: dotnet-hosting-6.0.30-win.exe windowsdesktop-runtime-6.0.30-win-x64.exe | .NET 6.0.30 Security Update for x64 Server | .NET 6.0 |
| KB5038283 | 17-Jun-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 17-Jun-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – x5.124 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| Microsoft Edge | 17-Jun-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 125 Update for x64 based Editions (Build 125.0.2535.51) Microsoft Edge-WebView2 Runtime Version 125 Update for x64 based Editions (Build 125.0.2535.51) | Microsoft Edge (Build 125.0.2535.51) | Microsoft Edge |
| KB5036896 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5037336 | 10-May-2024 | Solution Available | Install: dotnet-hosting-6.0.29-win-x64.exe windowsdesktop-runtime-6.0.29-win-x64.exe | .NET 6.0.29 Security Update for x64 Server | .NET 6.0 |
| KB890830* | 10-May-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.123 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5037034 | 10-May-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 10-May-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| Microsoft Edge | 10-May-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 123 Update for x64 based Editions (Build 123.0.2420.65) Microsoft Edge-WebView2 Runtime Version 123 Update for x64 based Editions (Build 123.0.2420.65) | Microsoft Edge (Build 123.0.2420.65) | Microsoft Edge |
| Microsoft Edge | 22-Apr-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 123 Update for x64 based Editions (Build 123.0.2420.65) Microsoft Edge-WebView2 Runtime Version 123 Update for x64 based Editions (Build 123.0.2420.65) | Microsoft Edge (Build 123.0.2420.65) | Microsoft Edge |
| Microsoft Edge | 20-Apr-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 122 Update for x64 based Editions (Build 122.0.2365.52) Microsoft Edge-WebView2 Runtime Version 122 Update for x64 based Editions (Build 122.0.2365.52) | Microsoft Edge (Build 122.0.2365.52) | Microsoft Edge |
| KB4052623 | 22-Apr-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Apr-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.122 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5036450 | 22-Apr-2024 | Solution Available | Install: dotnet-hosting-6.0.28-win.exe windowsdesktop-runtime-6.0.28-win-x64.exe | .NET 6.0.28 Security Update for x64 Server | .NET 6.0 |
| KB5035849 | 22-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| 7-Zip | 01-Apr-2024 | Solution Available | Install Recommended Update | 7-Zip (Build 23.01) | 7-Zip |
| KB5035119 | 01-Apr-2024 | Solution Available | Install: dotnet-hosting-6.0.27-win.exe windowsdesktop-runtime-6.0.27-win-x64.exe | .NET 6.0.27 Security Update for x64 Server | .NET 6.0 |
| KB890830* | 01-Apr-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.121 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5034683 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5034768 | 01-Apr-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5031908 | 01-Apr-2024 | Solution Available | Install Recommended Update | SQL Server 2019 RTM Cumulative Update (CU) 24 | Microsoft SQL Server 2019 |
| Microsoft Edge | 22-Feb-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 121 Update for x64 based Editions (Build 121.0.2277.83) Microsoft Edge-WebView2 Runtime Version 121 Update for x64 based Editions (Build 121.0.2277.83) | Microsoft Edge (Build 121.0.2277.83) | Microsoft Edge |
| KB890830* | 22-Feb-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.120 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5033733 | 22-Feb-2024 | Solution Available | Install: dotnet-hosting-6.0.26-win.exe windowsdesktop-runtime-6.0.26-win-x64.exe | .NET 6.0.26 Security Update for x64 Server | .NET 6.0 |
| KB5034273 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5034127 | 22-Feb-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5030333 | 25-Jan-2024 | Solution Available | Install Recommended Update | SQL Server 2019 RTM Cumulative Update (CU) 23 | Microsoft SQL Server 2019 |
| KB5033371 | 25-Jan-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 25-Jan-2024 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 21-Dec-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.119 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 21-Dec-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5032337 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5032196 | 21-Dec-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5032883 | 21-Dec-2023 | Solution Available | Install: dotnet-hosting-6.0.25-win.exe windowsdesktop-runtime-6.0.25-win-x64.exe | .NET 6.0.25 Security Update for x64 Server | .NET 6.0 |
| KB890830* | 29-Nov-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.118 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5032874 | 29-Nov-2023 | Solution Available | Install: dotnet-hosting-6.0.24-win.exe windowsdesktop-runtime-6.0.24-win-x64.exe | .NET 6.0.24 Security Update for x64 Server | .NET 6.0 |
| KB5031361 | 29-Nov-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 29-Nov-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 26-Oct-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.117 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5030559 | 26-Oct-2023 | Solution Available | Install: dotnet-hosting-6.0.22-win windowsdesktop-runtime-6.0.22-win-x64 | .NET 6.0.22 Security Update for x64 Server | .NET 6.0 |
| KB5021953 | 26-Oct-2023 | Solution Available | Install: dotnet-hosting-3.1.32-win | .NET Core 3.1.32 Security Update for x64 Server | .NET Core 3.1 |
| KB5030214 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5030178 | 26-Oct-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 26-Oct-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 14-Sep-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.116 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5029247 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5029647 | 14-Sep-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB4052623 | 14-Sep-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5024276 | 28-Aug-2023 | Solution Available | Install Recommended Update | SQL Server 2019 RTM Cumulative Update (CU) 20 | Microsoft SQL Server 2019 |
| KB890830* | 28-Aug-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.115 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 28-Aug-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5028168 | 28-Aug-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 25-Jul-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.114 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 25-Jul-2023 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5027536 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5027222 | 25-Jul-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB890830* | 12-Jun-2023 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.113 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 12-Jun-2023 | Solution Available | Install Recommended Update | Update for Microsoft Defender Antivirus antimalware platform (Version 4.18.2304.8) | Microsoft Defender Antivirus (platform update) |
| KB5026362 | 12-Jun-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5025229 | 11-May-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5023702 | 12-Apr-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5022840 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5022782 | 21-Mar-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5022286 | 07-Feb-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5021237 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5021085 | 24-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5020685 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5019966 | 10-Jan-2023 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5018419 | 14-Dec-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5017315 | 27-Oct-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5016623 | 14-Sep-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5012170 | 14-Sep-2022 | Solution Available | Install Recommended Update | Security Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5015811 | 03-Aug-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5014692 | 25-Jul-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5013941 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5013868 | 13-Jun-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5012647 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5012328 | 04-May-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5011503 | 18-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5010351 | 01-Apr-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5009718 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 1809 Windows Server 2019 |
| KB5009557 | 25-Feb-2022 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 1809 Windows Server 2019 |
| KB5005112 | 25-Feb-2022 | Solution Available | Install Recommended Update | Service Stack Update | Windows 10 Version 1809 Windows Server 2019 |

## Windows 10 IoT Enterprise LTSC v21H2

| Release: | PIC iX 4.3.2 and later | OPERATING_SYSTEM / SERVICE_PACK: (The Main Operating System is the OS of the system component that connects to the network) | PIC iX Servers and Client Application: Windows 10 IoT Enterprise LTSC v21H2 |
| --- | --- | --- | --- |

| Windows 10 IoT Enterprise LTSC v21H2 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Vulnerability / Patch ID | Last Update by Philips | Activity Status | Recommended Customer Action | Notes / Instructions | Microsoft Product |
| KB5087544 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5088859 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 27-May-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.141 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5093447 | 27-May-2026 | Solution Available | Install: dotnet-hosting-8.0.27-win.exe windowsdesktop-runtime-8.0.27-win-x64.exe | .NET 8.0.27 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 27-May-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 148 Update for x64 based Editions (Build 148.0.3967.54) Microsoft Edge-WebView2 Runtime Version 148 Update for x64 based Editions (Build 148.0.3967.54) | Microsoft Edge (Build 148.0.3967.54) | Microsoft Edge |
| KB5082200 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5084067 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 21-Apr-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.140 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 21-Apr-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5086096 | 21-Apr-2026 | Solution Available | Install: dotnet-hosting-8.0.26-win.exe for x64 Client windowsdesktop-runtime-8.0.26-win-x64.exe | .NET 8.0.26 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 21-Apr-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 147 Update for x64 based Editions (Build 147.0.3912.60) Microsoft Edge-WebView2 Runtime Version 147 Update for x64 based Editions (Build 147.0.3912.60) | Microsoft Edge (Build 147.0.3912.60) | Microsoft Edge |
| KB5078885 | 20-Mar-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5081277 | 20-Mar-2026 | Solution Available | Install: dotnet-hosting-8.0.25-win.exe for x64 Client windowsdesktop-runtime-8.0.25-win-x64.exe | .NET 8.0.25 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 20-Mar-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 145 Update for x64 based Editions (Build 145.0.3800.97) Microsoft Edge-WebView2 Runtime Version 145 Update for x64 based Editions (Build 145.3800.97) | Microsoft Edge (Build 145.0.3800.97) | Microsoft Edge |
| 7-Zip | 20-Mar-2026 | Solution Available | Install Recommended Update | 7-Zip Version 26.00 | 7-Zip |
| CVE-2025-29803 | 20-Mar-2026 | Solution Available | Install: SQL Server Management Studio 20.2: Version 20.2.37.0 Visual Studio Tools for Applications (VSTA) 2019: Version 16.0.35907 | Visual Studio Tools for Applications and SQL Server Management Studio Elevation of Privilege Vulnerability | Windows 10 Version 21H2 |
| KB5078297 |  | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU23) | Microsoft SQL Server 2022 |
| KB5075912 | 23-Feb-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 23-Feb-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 23-Feb-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.139 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077863 | 23-Feb-2026 | Solution Available | Install: dotnet-hosting-8.0.24-win.exe windowsdesktop-runtime-8.0.24-win-x64.exe | .NET 8.0.24 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 23-Feb-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 144 Update for x64 based Editions (Build 144.0.3719.115) Microsoft Edge-WebView2 Runtime Version 144 Update for x64 based Editions (Build 144.0.3719.115) | Microsoft Edge (Build 144.0.3719.115) | Microsoft Edge |
| KB5073724 | 26-Jan-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5074752 | 26-Jan-2026 | Solution Available | Install: dotnet-hosting-8.0.23-win.exe windowsdesktop-runtime-8.0.23-win-x64.exe | .NET 8.0.23 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 26-Jan-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.139) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build ​143.0.3650.139) | Microsoft Edge (Build 143.0.3650.139) | Microsoft Edge |
| KB5071546 | 17-Dec-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 17-Dec-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 17-Dec-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.138 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| Microsoft Edge | 17-Dec-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.66) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build ​143.0.3650.66) | Microsoft Edge (Build 143.0.3650.66) | Microsoft Edge |
| KB5068450 | 17-Dec-2025 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU22) | Microsoft SQL Server 2022 |
| KB5068781 | 25-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB890830* | 25-Nov-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.137 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5071066 | 25-Nov-2025 | Solution Available | Install: dotnet-hosting-8.0.22-win.exe windowsdesktop-runtime-8.0.22-win-x64.exe | .NET 8.0.22 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 25-Nov-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 142 Update for x64 based Editions (Build 142.0.3595.69) Microsoft Edge-WebView2 Runtime Version 142 Update for x64 based Editions (Build 142.0.3595.69) | Microsoft Edge (Build 142.0.3595.69) | Microsoft Edge |
| KB5066791 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5066746 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB4052623 | 22-Oct-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Oct-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.136 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5068331 | 22-Oct-2025 | Solution Available | Install: dotnet-hosting-8.0.21-win.exe windowsdesktop-runtime-8.0.21-win-x64.exe | .NET 8.0.21 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 22-Oct-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 141 Update for x64 based Editions (Build 141.0.3537.71) Microsoft Edge-WebView2 Runtime Version 141 Update for x64 based Editions (Build 141.0.3537.71) | Microsoft Edge (Build 141.0.3537.71) | Microsoft Edge |
| 7-Zip | 22-Oct-2025 | Solution Available | Install Recommended Update | 7-Zip (Build 25.01) | 7-Zip |
| KB5065429 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5065956 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB5066233 | 17-Sep-2025 | Solution Available | Install: dotnet-hosting-8.0.20-win.exe windowsdesktop-runtime-8.0.20-win-x64.exe | .NET 8.0.20 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 17-Sep-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 140 Update for x64 based Editions (Build 140.0.3485.54) Microsoft Edge-WebView2 Runtime Version 140 Update for x64 based Editions (Build 140.0.3485.54) | Microsoft Edge (Build 140.0.3485.54) | Microsoft Edge |
| KB5063709 | 20-Aug-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB890830* | 20-Aug-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.1​35 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 20-Aug-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5064838 | 20-Aug-2025 | Solution Available | Install: dotnet-hosting-8.0.19-win.exe windowsdesktop-runtime-8.0.19-win-x64.exe | .NET 8.0.19 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 20-Aug-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 139 Update for x64 based Editions (Build 139.0.3405.86) Microsoft Edge-WebView2 Runtime Version 139 Update for x64 based Editions (Build 139.0.3405.86) | Microsoft Edge (Build 139.0.3405.86) | Microsoft Edge |
| KB5059390 | 20-Aug-2025 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU20) | Microsoft SQL Server 2022 |
| KB5062554 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5062153 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB5063326 | 20-Jul-2025 | Solution Available | Install: dotnet-hosting-8.0.18-win.exe windowsdesktop-runtime-8.0.18-win-x64.exe | .NET 8.0.18 Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 20-Jul-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 138 Update for x64 based Editions (Build 138.0.3351.77) Microsoft Edge-WebView2 Runtime Version 138 Update for x64 based Editions (Build 138.0.3351.77) | Microsoft Edge (Build 138.0.3351.77) | Microsoft Edge |
| KB5060533 | 19-Jun-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 19-Jun-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 19-Jun-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – x5.134 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5061935 | 19-Jun-2025 | Solution Available | Install: Dotnet-hosting-8.0.17-win.exe Windowsdesktop-runtime-8.0.17-win-x64.exe | .NET 8.0.17 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 19-Jun-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 137 Update for x64 based Editions (Build 137.0.3296.68) Microsoft Edge-WebView2 Runtime Version 137 Update for x64 based Editions (Build 137.0.3296.68) | Microsoft Edge (Build 137.0.3296.68) | Microsoft Edge |
| KB5058379 | 22-May-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 22-May-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-May-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.133 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5059200 | 22-May-2025 | Solution Available | Install: dotnet-hosting-8.0.16-win.exe windowsdesktop-runtime-8.0.16-win-x64.exe | .NET 8.0.16 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 22-May-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 136 Update for x64 based Editions (Build 136.0.3240.64) Microsoft Edge-WebView2 Runtime Version 136 Update for x64 based Editions (Build 136.0.3240.64) | Microsoft Edge (Build 136.0.3240.64) | Microsoft Edge |
| KB5050771 | 22-May-2025 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU18) | Microsoft SQL Server 2022 |
| KB5055518 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5055682 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB5056686 | 24-Apr-2025 | Solution Available | Install: Dotnet-hosting-8.0.15-win.exe Windowsdesktop-runtime-8.0.15-win-x64.exe | .NET 8.0.15 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 24-Apr-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 135 Update for x64 based Editions (Build 135.0.3179.54) Microsoft Edge-WebView2 Runtime Version 135 Update for x64 based Editions (Build 135.0.3179.54) | Microsoft Edge (Build 135.0.3179.54) | Microsoft Edge |
| KB5053606 | 28-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 28-Mar-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5054229 | 28-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.14-win.exe windowsdesktop-runtime-8.0.14-win-x64.exe | .NET 8.0.14 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 28-Mar-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 134 Update for x64 based Editions (Build 134.0.3124.51) Microsoft Edge-WebView2 Runtime Version 134 Update for x64 based Editions (Build 134.0.3124.51) | Microsoft Edge (Build 134.0.3124.51) | Microsoft Edge |
| KB5051974 | 03-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB890830* | 03-Mar-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.132 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5052978 | 03-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.13-win.exe windowsdesktop-runtime-8.0.13-win-x64.exe | .NET 8.0.13 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 03-Mar-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 133 Update for x64 based Editions (Build 133.0.3065.59) Microsoft Edge-WebView2 Runtime Version 133 Update for x64 based Editions (Build 133.0.3065.59) | Microsoft Edge (Build 133.0.3065.59) | Microsoft Edge |
| KB5049981 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5050416 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 03-Feb-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.131 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5050525 | 03-Feb-2025 | Solution Available | Install: Dotnet-hosting-8.0.12-win.exe Windowsdesktop-runtime-8.0.12-win-x64.exe | .NET 8.0.12 Security Update for x64 Client | .NET 8.0 |
| Microsoft Edge | 03-Feb-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 132 Update for x64 based Editions (Build 132.0.2957.115) Microsoft Edge-WebView2 Runtime Version 132 Update for x64 based Editions (Build 132.0.2957.115) | Microsoft Edge (Build 132.0.2957.115) | Microsoft Edge |
| KB5048652 | 19-Dec-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| 7-Zip | 19-Dec-2024 | Solution Available | Install Recommended Update | 7-Zip (Build 24.09) | 7-Zip |
| KB5048033 | 19-Dec-2024 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU) 16 | Microsoft SQL Server 2022 |
| KB5046613 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5046541 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 25-Nov-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.130 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5047489 | 25-Nov-2024 | Solution Available | Install: dotnet-hosting-8.0.11-win.exe windowsdesktop-runtime-8.0.11-win-x64.exe | .NET 8.0.11 Security Update for x64 Client (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 25-Nov-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 131 Update for x64 based Editions (Build 131.0.2903.48) Microsoft Edge-WebView2 Runtime Version 131 Update for x64 based Editions (Build 131.0.2903.48) | Microsoft Edge (Build 131.0.2903.48) | Microsoft Edge |

## Windows Server 2022

| Release: | PIC iX 4.3.2 and later | OPERATING_SYSTEM / SERVICE_PACK: (The Main Operating System is the OS of the system component that connects to the network) | PIC iX Servers and Client Application: Windows Server 2022 |
| --- | --- | --- | --- |

| Windows Server 2022 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Vulnerability / Patch ID | Last Update by Philips | Activity Status | Recommended Customer Action | Notes / Instructions | Microsoft Product |
| KB5087545 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5088862 | 27-May-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 27-May-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.141 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5093447 | 27-May-2026 | Solution Available | Install: dotnet-hosting-8.0.27-win.exe windowsdesktop-runtime-8.0.27-win-x64.exe | .NET 8.0.27 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 27-May-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 148 Update for x64 based Editions (Build 148.0.3967.54) Microsoft Edge-WebView2 Runtime Version 148 Update for x64 based Editions (Build 148.0.3967.54) | Microsoft Edge (Build 148.0.3967.54) | Microsoft Edge |
| KB5082142 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5084071 | 21-Apr-2026 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 21-Apr-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.140 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 21-Apr-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5086096 | 21-Apr-2026 | Solution Available | Install: dotnet-hosting-8.0.26-win.exe for x64 Client windowsdesktop-runtime-8.0.26-win-x64.exe | .NET 8.0.26 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 21-Apr-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 147 Update for x64 based Editions (Build 147.0.3912.60) Microsoft Edge-WebView2 Runtime Version 147 Update for x64 based Editions (Build 147.0.3912.60) | Microsoft Edge (Build 147.0.3912.60) | Microsoft Edge |
| KB5078766 | 20-Mar-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5081277 | 20-Mar-2026 | Solution Available | Install: dotnet-hosting-8.0.25-win.exe for x64 Client windowsdesktop-runtime-8.0.25-win-x64.exe | .NET 8.0.25 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 20-Mar-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 145 Update for x64 based Editions (Build 145.0.3800.97) Microsoft Edge-WebView2 Runtime Version 145 Update for x64 based Editions (Build 145.3800.97) | Microsoft Edge (Build 145.0.3800.97) | Microsoft Edge |
| 7-Zip | 20-Mar-2026 | Solution Available | Install Recommended Update | 7-Zip Version 26.00 | 7-Zip |
| CVE-2025-29803 | 20-Mar-2026 | Solution Available | Install: SQL Server Management Studio 20.2: Version 20.2.37.0 Visual Studio Tools for Applications (VSTA) 2019: Version 16.0.35907 | Visual Studio Tools for Applications and SQL Server Management Studio Elevation of Privilege Vulnerability | Windows 10 Version 21H2 |
| KB5078297 | 20-Mar-2026 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU23) | Microsoft SQL Server 2022 |
| KB5075906 | 23-Feb-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 23-Feb-2026 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 23-Feb-2026 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.139 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5077863 | 23-Feb-2026 | Solution Available | Install: dotnet-hosting-8.0.24-win.exe windowsdesktop-runtime-8.0.24-win-x64.exe | .NET 8.0.24 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 23-Feb-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 144 Update for x64 based Editions (Build 144.0.3719.115) Microsoft Edge-WebView2 Runtime Version 144 Update for x64 based Editions (Build 144.0.3719.115) | Microsoft Edge (Build 144.0.3719.115) | Microsoft Edge |
| KB5073457 | 26-Jan-2026 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5074752 | 26-Jan-2026 | Solution Available | Install: dotnet-hosting-8.0.23-win.exe windowsdesktop-runtime-8.0.23-win-x64.exe | .NET 8.0.23 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 26-Jan-2026 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.139) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build ​143.0.3650.139) | Microsoft Edge (Build 143.0.3650.139) | Microsoft Edge |
| KB5071547 | 17-Dec-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 17-Dec-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 17-Dec-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.138 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| Microsoft Edge | 17-Dec-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 143 Update for x64 based Editions (Build 143.0.3650.66) Microsoft Edge-WebView2 Runtime Version 143 Update for x64 based Editions (Build ​143.0.3650.66) | Microsoft Edge (Build 143.0.3650.66) | Microsoft Edge |
| KB5068450 | 17-Dec-2025 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU22) | Microsoft SQL Server 2022 |
| KB5068787 | 25-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB890830* | 25-Nov-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.137 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5071066 | 25-Nov-2025 | Solution Available | Install: dotnet-hosting-8.0.22-win.exe windowsdesktop-runtime-8.0.22-win-x64.exe | .NET 8.0.22 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 25-Nov-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 142 Update for x64 based Editions (Build 142.0.3595.69) Microsoft Edge-WebView2 Runtime Version 142 Update for x64 based Editions (Build 142.0.3595.69) | Microsoft Edge (Build 142.0.3595.69) | Microsoft Edge |
| KB5070884 | 12-Nov-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5066782 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for Microsoft | Windows 10 Version 21H2 |
| KB5066743 | 22-Oct-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB4052623 | 22-Oct-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-Oct-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.136 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5068331 | 22-Oct-2025 | Solution Available | Install: dotnet-hosting-8.0.21-win.exe windowsdesktop-runtime-8.0.21-win-x64.exe | .NET 8.0.21 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 22-Oct-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 141 Update for x64 based Editions (Build 141.0.3537.71) Microsoft Edge-WebView2 Runtime Version 141 Update for x64 based Editions (Build 141.0.3537.71) | Microsoft Edge (Build 141.0.3537.71) | Microsoft Edge |
| 7-Zip | 22-Oct-2025 | Solution Available | Install Recommended Update | 7-Zip (Build 25.01) | 7-Zip |
| KB5065432 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for Microsoft | Windows 10 Version 21H2 |
| KB5065962 | 17-Sep-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB5066233 | 17-Sep-2025 | Solution Available | Install: dotnet-hosting-8.0.20-win.exe windowsdesktop-runtime-8.0.20-win-x64.exe | .NET 8.0.20 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 17-Sep-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 140 Update for x64 based Editions (Build 140.0.3485.54) Microsoft Edge-WebView2 Runtime Version 140 Update for x64 based Editions (Build 140.0.3485.54) | Microsoft Edge (Build 140.0.3485.54) | Microsoft Edge |
| KB5063880 | 20-Aug-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB890830* | 20-Aug-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 - v5.1​35 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB4052623 | 20-Aug-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5064838 | 20-Aug-2025 | Solution Available | Install: dotnet-hosting-8.0.19-win.exe windowsdesktop-runtime-8.0.19-win-x64.exe | .NET 8.0.19 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 20-Aug-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 139 Update for x64 based Editions (Build 139.0.3405.86) Microsoft Edge-WebView2 Runtime Version 139 Update for x64 based Editions (Build 139.0.3405.86) | Microsoft Edge (Build 139.0.3405.86) | Microsoft Edge |
| KB5059390 | 20-Aug-2025 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU20) | Microsoft SQL Server 2022 |
| KB5062572 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5062159 | 20-Jul-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB5063326 | 20-Jul-2025 | Solution Available | Install: dotnet-hosting-8.0.18-win.exe windowsdesktop-runtime-8.0.18-win-x64.exe | .NET 8.0.18 Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 20-Jul-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 138 Update for x64 based Editions (Build 138.0.3351.77) Microsoft Edge-WebView2 Runtime Version 138 Update for x64 based Editions (Build 138.0.3351.77) | Microsoft Edge (Build 138.0.3351.77) | Microsoft Edge |
| KB5060526 | 19-Jun-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 19-Jun-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus | Microsoft Defender Antivirus (platform update) |
| KB890830* | 19-Jun-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.134 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5061935 | 19-Jun-2025 | Solution Available | Install: dotnet-hosting-8.0.17-win-x64.exe windowsdesktop-runtime-8.0.17-win-x64.exe | .NET 8.0.17 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 19-Jun-2025 | Solution Available | Install: Microsoft Edge-Stable Version 137 Update for x64 based Editions (Build 137.0.3296.68) Microsoft Edge-WebView2 Runtime Version 137 Update for x64 based Editions (Build 137.0.3296.68) | Microsoft Edge (Build 137.0.3296.68) | Microsoft Edge |
| KB5058385 | 22-May-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 22-May-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB890830* | 22-May-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.133 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5059200 | 22-May-2025 | Solution Available | Install: dotnet-hosting-8.0.16-win.exe windowsdesktop-runtime-8.0.16-win-x64.exe | .NET 8.0.16 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 22-May-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 136 Update for x64 based Editions (Build 136.0.3240.64) Microsoft Edge-WebView2 Runtime Version 136 Update for x64 based Editions (Build 136.0.3240.64) | Microsoft Edge (Build 136.0.3240.64) | Microsoft Edge |
| KB5050771 | 22-May-2025 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU18) | Microsoft SQL Server 2022 |
| KB5055526 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5055688 | 24-Apr-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB5056686 | 24-Apr-2025 | Solution Available | Install: dotnet-hosting-8.0.15-win.exe windowsdesktop-runtime-8.0.15-win-x64.exe | .NET 8.0.15 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 24-Apr-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 135 Update for x64 based Editions (Build 135.0.3179.54) Microsoft Edge-WebView2 Runtime Version 135 Update for x64 based Editions (Build 135.0.3179.54) | Microsoft Edge (build 135.0.3179.54) | Microsoft Edge |
| KB5053603 | 28-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB4052623 | 28-Mar-2025 | Solution Available | Update to latest definition files per Antivirus guide | Update for Microsoft Defender Antivirus antimalware platform | Microsoft Defender Antivirus (platform update) |
| KB5054229 | 28-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.14-win.exe windowsdesktop-runtime-8.0.14-win-x64.exe | .NET 8.0.14 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 28-Mar-2025 | Solution Available | Install: Microsoft Edge-Extended Stable Channel Version 134 Update for x64 based Editions (Build 134.0.3124.51) Microsoft Edge-WebView2 Runtime Version 134 Update for x64 based Editions (Build 134.0.3124.51) | Microsoft Edge (Build 134.0.3124.51) | Microsoft Edge |
| KB5051979 | 03-Mar-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB890830* | 03-Mar-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.132 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5052978 | 03-Mar-2025 | Solution Available | Install: dotnet-hosting-8.0.13-win.exe windowsdesktop-runtime-8.0.13-win-x64.exe | .NET 8.0.13 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 03-Mar-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 133 Update for x64 based Editions (Build 133.0.3065.59) Microsoft Edge-WebView2 Runtime Version 133 Update for x64 based Editions (Build 133.0.3065.59) | Microsoft Edge (Build 133.0.3065.59) | Microsoft Edge |
| KB5049983 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5050187 | 03-Feb-2025 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 03-Feb-2025 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.131 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5050525 | 03-Feb-2025 | Solution Available | Install: Dotnet-hosting-8.0.12-win.exe Windowsdesktop-runtime-8.0.12-win-x64.exe | .NET 8.0.12 Security Update for x64 Server | .NET 8.0 |
| Microsoft Edge | 03-Feb-2025 | Solution Available | Install: Microsoft Edge-Stable Channel Version 132 Update for x64 based Editions (Build 132.0.2957.115) Microsoft Edge-WebView2 Runtime Version 132 Update for x64 based Editions (Build 132.0.2957.115) | Microsoft Edge (Build 132.0.2957.115) | Microsoft Edge |
| KB5048654 | 19-Dec-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| 7-Zip | 19-Dec-2024 | Solution Available | Install Recommended Update | 7-Zip (Build 24.09) | 7-Zip |
| KB5048033 | 19-Dec-2024 | Solution Available | Install Recommended Update | SQL Server 2022 RTM Cumulative Update (CU) 16 | Microsoft SQL Server 2022 |
| KB5046616 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for Windows | Windows 10 Version 21H2 |
| KB5046547 | 25-Nov-2024 | Solution Available | Install Recommended Update | Cumulative Update for .NET Framework | Windows 10 Version 21H2 |
| KB890830* | 25-Nov-2024 | Solution Available | Install Recommended Update | Windows Malicious Software Removal Tool x64 – v5.130 | Microsoft Windows Malicious Software Removal Tool (MSRT) |
| KB5047489 | 25-Nov-2024 | Solution Available | Install: dotnet-hosting-8.0.11-win.exe windowsdesktop-runtime-8.0.11-win-x64.exe | .NET 8.0.11 Security Update for x64 Server (Applicable for PIC iX 4.3) | .NET 8.0 |
| Microsoft Edge | 25-Nov-2024 | Solution Available | Install: Microsoft Edge-Stable Channel Version 131 Update for x64 based Editions (Build 131.0.2903.48) Microsoft Edge-WebView2 Runtime Version 131 Update for x64 based Editions (Build 131.0.2903.48) | Microsoft Edge (Build 131.0.2903.48) | Microsoft Edge |

## Validated AVS

| Validated AVS: | PIC iX C.03; PIC iX 4.x. McAfee Virus Scan Enterprise 8.7i; No; No. McAfee Virus Scan Enterprise 8.8; Yes, requires minimum Patch 9; No. McAfee Endpoint Security 10.6.1; Yes; Yes (4.0 & 4.1) No (4.2 & 4.3). McAfee Endpoint Security 10.7; Yes; Yes (4.0 & 4.1) No (4.2 & 4.3). Trellix (formerly McAfee) Endpoint Security 10.7; No; Yes (4.x - Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). Symantec Endpoint Protection 11.0.4; No; No. Symantec Endpoint Protection 12.1; No; No. Symantec Endpoint Protection 12.1.6; Yes; No. Symantec Endpoint Protection 14.2; Yes; Yes (4.0 & 4.1 & 4.2) No (4.3). Symantec Endpoint Protection 14.3 RU6; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 & 4.4). Symantec Endpoint Protection 14.3 RU9; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 & 4.4 & 4.5). Symantec Endpoint Protection 14.3 RU10; Yes; Yes (4.x - Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). Trend Micro Deep Security 9.6.2; No; No. Trend Micro Deep Security 10.0; Yes; No. Trend Micro Deep Security 11.0; Yes; Yes (4.0) No (4.1 & 4.2 & 4.3). Trend Micro Deep Security 12.0; Yes; Yes (4.0) No (4.1 & 4.2 & 4.3). Trend Micro Deep Security 20.0; Yes; No. Trend Micro Deep Security 20.0.1; No; Yes (4.3.2 and up & 4.4 & 4.5 – Windows 10 IoT Enterprise LTSC v21H2 and Windows Server 2022). Trend Micro Deep Security 20.0.2; No; Yes (4.4 and greater – Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.26; Yes; No. CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.33; Yes; No. CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.37; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 & 4.4). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.40; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 & 4.4). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.46; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 & 4.4). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.50; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 & 4.4). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.54; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.56; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 6.58; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.06; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.14; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.16; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 & 4.4). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.19; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 – Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.21; Yes; Yes (4.0 & 4.1 & 4.2 & 4.3 – Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.23; Yes; Yes (4.3 & 4.4 & 4.5 – Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.25; Yes; Yes (4.4 & 4.5 – Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.28; Yes; Yes (4.4 – Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). CrowdStrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.32; Yes; Yes (4.x – Windows 10 Enterprise LTSC v1809 and v21H2/Windows Server 2019/Windows Server 2022). Crowdstrike Falcon Prevent Next Generation Antivirus, Falcon Sensor for Windows 7.34; No; Yes (4.x – Windows 10 Enterprise LTSC v1809, 21H2, Windows Server 2019, Windows Server 2022. Windows Defender for Windows 10 v1607 and Windows Server 2016; Yes; No. Windows Defender for Windows 10 v1809 and Windows Server 2019; No; Yes (4.x). Windows Defender for Windows 10 v21H2 and Windows Server 2022; No; Yes (4.x). Microsoft Defender for Endpoint Protection (managed via MECM EIC agent only); No; Yes (4.3.2 and greater – Windows v21H2 and Windows Server 2022 only). Refer to the effective Antivirus Software Use and Configuration Guide for vendor specific installation and configuration instructions for your version of PIC iX. |
| --- | --- |
| Network ports for clinical use: | Refer to Philips Patient Monitoring System Security for Clinical Networks |
| Remote Service Connectivity: | Refer to Philips Patient Monitoring System Security for Clinical Networks |

## Hypervisor Compatibility

| Platform | PIC iX C.03.xx | PIC iX 4.x |
| --- | --- | --- |
| VMware ESXi 4.0, 4.1 | NO | NO |
| VMware (vSphere) 5.0 U3, 5.1 and 5.5 (ESXi and vCenter Server) | NO | NO |
| VMware ESXi 6.0 | YES | NO |
| VMware ESXi 6.5 | YES | NO |
| VMware ESXi 6.7 | YES | YES (4.0) NO (4.1 & 4.2 & 4.3) |
| VMware ESXi 7.0 | YES | YES (4.0 & 4.1 & 4.2 & 4.3 & 4.4 & 4.5) |
| VMware ESXi 8.0 (Hardware version 17-21) | NO | YES (4.3.2 and greater – Windows Server 2022) |
| VMware ESXi 9.0 | NO | YES (4.3.2 and greater – Windows 10 v1809, 21H2, Windows Server 2019, 2022) |
| Hyper-V Server 2008 R2 and 2012 | NO | NO |
| Hyper-V Server 2012 R2 | NO | NO |
| Hyper-V Server 2016 (VHDX ONLY) | YES | YES (4.0 & 4.1) NO (4.3) |
| Hyper-V Server 2019 (VHDX ONLY) | YES | YES (4.0 & 4.1 & 4.2 & 4.3) |
| Hyper-V Server 2022 (VHDX ONLY) | NO | YES (4.3.2 and greater – Windows Server 2022) |
| Hyper-V Server 2025 (VHDX ONLY) | NO | YES (4.3.2 and greater – Windows Server 2022) |
| Nutanix AOS 5.15 (LTS) with AHV | NO | YES (4.0) |
| Nutanix AOS 5.20 (LTS) with AHV | YES | YES (4.0 & 4.1) |
| Nutanix AOS 6.5 (LTS) with AHV | YES | YES (4.2 & 4.3 & 4.4 & 4.5) |
| Nutanix AOS 6.10 (LTS) with AHV | YES | YES (4.2 & 4.3 & 4.6 & 4.7) |
| Nutanix AOS 7.3 (LTS) with AHV | NO | YES (4.3.2 and greater – Windows Server 2022) |

Incompatible patches

This table lists the patches that have been found incompatible and should not be installed on the product. It is strongly advised to follow the recommended customer action.

| Incompatible patch | Evaluation date | Applicable for release | Recommended customer action |
| --- | --- | --- | --- |
| N/A | N/A | N/A | N/A |

Evaluation date:

The date shown in column “Evaluation date” defines the last date that the status has been (re-)assessed by Philips.

Scoring:

This risk rating is a risk indication completed objectively by Philips, based on the product’s intended use and intended operating environment as specified in the Instructions for use (IFU) and other product documentation. Rating includes mitigating factors designed into the product (as for example documented in the MDS2) and in the operating environment. The Customer will have to assess whether the risk indication may differ for its own organization. The < Scoring System, for example CVSS 3.1> scoring needs to be adjusted by the customer if the device is used differently from the intended use and intended operating environment. The Customer is solely responsible for pursuing alternative measures if the Customer considers the risk indication to be higher than the Risk Indication contained in this report.

Applicable for release:

This column holds the versions (Software releases) to which the entry applies.

Recommended Customer Action:

Refer to tables above for recommended customer actions.